|  |
| --- |
| STANDARD OPERATING PROCEDURE |
| |  |  | | --- | --- | | **Title: Liquid Chromatography using Accela system for MRM assays** | | | **Version #: 1.1** | **Author: Hui Zhang Laboratory – Johns Hopkins University** | | **Date: 06/10/2016** |  | |

# Purpose

The purpose of this document is to describe the liquid chromatography (LC) method for quantitative mass spectrometry-based analyses.

# Scope

This procedure describes the setup of the LC and the method parameters. It is specific to the operation of the Thermo Scientific Accela system.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures. The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

* HPLC: Accela 1250 Quaternary Low Pump (Thermo Scientific)
* Autosampler: Accela Open Autosampler (Thermo Scientific)

# Materials

* Injection loop: 10 µL
* Column heater: Hot pocket (Thermo Scientific; cat. # 92016-150)
* Analytical Column: 1.0 mm I.D. x 15 cm Viva C18 3 µm 300 Å (Restek; cat. # 9514361)
* Autosampler vials: Polypropylene 12 x 32 mm screw neck vial with caps and PTFE/silicone septum, 300 µL (Waters; cat. # 186002640)
* Water: Optima LC/MS-grade (Fisher Scientific; cat. # W6-4)
* Acetonitrile: Optima LC/MS-grade (Fisher Scientific; cat. # A955-4)
* Formic Acid: LC-MS Ultra (Sigma-Aldrich; cat. # 14265)

# Solutions

* Gradient pump, mobile phase A: Water/0.2 % formic acid
* Gradient pump, mobile phase B: 90% ACN/0.2% formic acid

# Procedure

1. Autosampler method

Temperature: 4°C

Syringe: 100 µL

Airgap Volume (µl): 3

Front Volume (µl): 2

Rear Volume (µl): 2

Filling Speed (µl/s): 5

Pullup Delay (ms): 3000

Inject to: LC Vlv1

Injection Speed (µl/s): 5

Pre & Post Inject Delay (ms): 500, 500

Needle Gap Valve Clean (mm): 3

Valve Clean Time Solvent 2 (s): 2

Post Clean Time Solvent 2 (s): 2

Valve Clean Time Solvent 1 (s): 5

Post Clean Time Solvent 1 (s): 5

Stator Wash: 0

Delay Stator Wash (s): 120

Stator Wash Time Solvent 2 (s): 5

Stator Wash Time Solvent 1 (s): 5

1. Gradient method:
   1. Flow rate: 50 µL/min
   2. Column heater temperature: 50 °C
   3. Timetable:

|  |  |  |
| --- | --- | --- |
| Retention (min) | Flow (µL/min) | %B |
| 0.00 | 50 | 1.00 |
| 5.00 | 50 | 1.00 |
| 30.00 | 50 | 35.00 |
| 35.00 | 50 | 55.00 |
| 38.00 | 50 | 95.00 |
| 45.00 | 50 | 95.00 |
| 46.00 | 50 | 1.00 |
| 60.00 | 50 | 1.00 |